## Instructions
Cell Color Indicator
Input | Output
Ellipsometric Parameters
Ψ | Δ
The overall goal of this Excel calculator is to allow one to experimentally determine thin film (on substrate) properties using a simple monochromatic rotating analyzer ellipsometer (RAE). This is done through regressive fitting of an optical model to experimental data using the Solver add-on in Excel.
Experimental Data
The "Exp. Data Collection" page is where one inputs the experimental intensity measurements made at various instrument angles of incidence and analyzer azimuths. The polarizer angle (α1) relative to the incident plane is typically fixed at 45 degrees. The ellipsometric parameters are immediately calculated and displayed. This page has its values referenced by "Exp. Data Processing".
The "Exp. Data Processing" page acts as a home page that summarizes the experimental data and its relationship to the fitted model.
Model Fitting
The regressive fitting process requires Excel's Solver add-on be enabled in the settings. It also requires that one knows the complex indicies of refraction for the sample ambient phase (likely air) and the substrate that the film is on. If the default solver settings are not automatically present they can be loaded using the "Default Solver Settings" cells below the parameter calculator table.
Regressive fitting involves varying the model ellipsometric parameters as a function of film refractive index and thickness. These values are varied until the average sum of normalized squared errors between model and experimental parameters is minimized. Solver ranges for the film refractive index (n1), coefficient of extinction (k1), and thickness (d1) can be used to decrease solving time and increase accuracy. Setting the "min" and "max" cells to the same value is equivalent to telling the solver that a parameter is known. This is particularly useful for the coefficient of extinction (k1) because it is typically considered or assumed nonexistent (k1 = 0) for many films of interest
RAE Instrument Configuration
Fundamental Ellipsometric Equation
𝛲=𝛲_π/𝛲_σ =tan⁡(ψ) 𝑒^𝑗Δ
Fundamental Ellipsometric Equation
Three-Phase Reflections
Experimental Data Processing

## Exp. Data Collection
Polarizer azimuth angle (α1) [°]
45
AoI [°] |  | INPUT |  |  |  | OUTPUT
[6 empty] | Ψ | Δ
20 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | #N/A | #N/A
Intensity
25 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 46.5798 | 170.754
| Intensity | 1 | 60 | 137 | 67
30 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 43.028 | 166.6645
| Intensity | 3 | 101 | 202 | 88
35 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 38.8613 | 164.7469
| Intensity | 11 | 211 | 373 | 137
40 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 35.6943 | 169.1093
| Intensity | 18 | 310 | 501 | 160
45 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 33.2109 | 172.6212
| Intensity | 25 | 357 | 524 | 153
50 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 32.2015 | 150.2136
| Intensity | 49 | 295 | 402 | 117
55 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 41.4376 | 129.2823
| Intensity | 50 | 145 | 219 | 113
60 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 64.3964 | 147.7017
| Intensity | 29 | 31 | 141 | 135
65 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | 39.6924 | 97.2101
| Intensity | 167 | 225 | 214 | 155
70 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | #N/A | #N/A
Intensity
75 | Analyzer α2 [°] | 45 | 90 | -45 | 0 | #N/A | #N/A
Intensity
ψ=atan⁡(tan⁡(α_1 )∗tan⁡(├ acos⁡((I_((90°) )−I_((0°) ))/(I_((90°) )+I_((0°) ) )) ┤/2) )
Δ=acos⁡(├ (I_((45°) )−I_((−45) ))/(I_((45°) )+I_((−45°) ) )┤/sin⁡(2 atan⁡(tan⁡(ψ)/tan⁡(α_1 ) ) ) )
Input relative light intensity measurements as a function of angle of incidence and analyzer azimuth angle. Ellipsometric paramteres Ψ and Δ are automatically calculated as a result. Measuring ellipsometric parameters at additional angles of incidence and analyzer azimuth will provide more data points for which to increase confidence that the fit model accurately describes the properties of the sample.

## Exp. Data Processing
Angle of | Experimental Ellipsometric Parameters |  |  | Model Ellipsometric Parameters |  |  | Fit % Error
Incidence (AoI)[o] | Ψ (Exp) [°] | Δ (Exp) [°] | P (Exp) | Ψ (Model) [°] | Δ (Model) [°] | P (Model) | Ψ (% Err) | Δ (% Err)
20 | #N/A | #N/A | #N/A | 45.2752 | 176.1033 | -1.00732029319911+0.0686143737503771i | #N/A | #N/A
25 | 46.5798 | 170.754 | -1.04299516908213+0.169787349169531i | 46.5329 | 176.4097 | -1.05292108795985+0.0660652073390131i | 0.1009% | 3.206%
30 | 43.028 | 166.6645 | -0.908258874667954+0.215297341599409i | 42.5934 | 168.1831 | -0.899850169319497+0.188266269322302i | 1.02% | 0.9029%
35 | 38.8613 | 164.7469 | -0.777399289099527+0.211989256410831i | 39.2383 | 168.1245 | -0.799215023327781+0.168063520618849i | 0.9608% | 2.009%
40 | 35.6943 | 169.1093 | -0.705482006339735+0.13573566586909i | 36.907 | 173.8999 | -0.746759322810276+0.0798066611936724i | 3.286% | 2.755%
45 | 33.2109 | 172.6212 | -0.649232370543845+0.0840759038574506i | 34.4253 | 172.7982 | -0.679957350090719+0.085920097970445i | 3.528% | 0.1024%
50 | 32.2015 | 150.2136 | -0.546566951031605+0.31284938473895i | 32.948 | 149.5372 | -0.558649923501517+0.328582148817699i | 2.266% | 0.4524%
55 | 41.4376 | 129.2823 | -0.558928342520188+0.683307729178604i | 42.4495 | 126.1505 | -0.539595320407904+0.738600806729938i | 2.384% | 2.483%
60 | 64.3964 | 147.7017 | -1.76394686907021+1.11504715271366i | 69.5371 | 147.821 | -2.26824980731351+1.4272328130306i | 7.393% | 0.08069%
65 | 39.6924 | 97.2101 | -0.104170312044327+0.823430285438589i | 38.1828 | 95.6566 | -0.0775163651498085+0.78260593305774i | 3.954% | 1.624%
70 | #N/A | #N/A | #N/A | 26.3514 | 87.4339 | 0.0221776820938759+0.494850610113642i | #N/A | #N/A
75 | #N/A | #N/A | #N/A | 17.3442 | 75.7302 | 0.0769812622306773+0.30267610120186i | #N/A | #N/A
Left graph displays the model fit to the experimental determinations of the ellipsometric parameters. Right graph displays error of ellipsometric parameters.

## Model Fitting
Index of Refraction Solver Inputs
Media | Symbol | Value | Units
0 - Atmosphere | n0 | 1
| k0 | 0
1 - Film | n1 | 1.483
| k1 | 0
| d1 | 14857 | Angs
2 - Substrate | n2 | 3.8435
| k2 | 0.015793
Ellipsometer Configuration:
Incident Wavelength | λ | 6530 | Angs
https://refractiveindex.info
Refractive Index (n1) Solver Range
Min | Max
1.483 | 1.483
Extinction Coefficient (k1) Solver Range
Min | Max
0 | 0
Thickness (t1) Solver Range [Angs]
Min | Max
5000 | 15000
Film 1 Solver Output
n1 | 1.483
k1 | 0
d1 [Angs] | 14857
Sum of Normalized Squared Errors
Ψ SSE Norm. | Δ SSE Norm. | Average SSE
0.010632 | 0.0031741 | 0.0069029
Solver Parameters | [5 empty] | Model Parameters
Ψ (Exp) [°] | Δ (Exp) [°] | Ψ (Model) [°] | Δ (Model) [°] | Ψ SE Normalized [°] | Δ SE Normalized [°] | 0 - Atmosphere |  | 1 - Film |  |  | 2 - Substrate |  | AoI | AoI | Wavelength | Media Complex Indicies of Refraction |  |  | Intermediate Parameters | [10 empty] | Ellipsometric Parameters
[6 empty] | n0 | k0 | n1 | k1 | d1 [Angs] | n2 | k2 | (Φ0) [°] | (Φ0) [Radians] | λ [Angs] | ñ0 | ñ1 | ñ2 | Cos(θ0) | Cos(θ1) | Cos(θ2) | ρ01,π | ρ01,σ | ρ12,π | ρ12,σ | β | Ρπ | Ρσ | P | Ψ [Radians] | Δ [Radians] | Ψ [°] | Δ [°]
|  | 44.9957 | 179.9909 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 1 | 0.017453 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.999847695156391 | 0.999930751020681 | 0.999989691223809+8.47197036533272E-08i | 0.194482793430097 | -0.19456271538857 | 0.443142762270463+0.00165101299261458i | -0.443190128992894-0.00165099482166483i | 21.1986832638836 | 0.223489446166602+0.424470296719345i | -0.223590488725572-0.42449934104805i | -0.999848294496414-0.000158121921026303i | 0.78532 | -3.1414 | 44.9957 | 179.9909
|  | 44.9826 | 179.9636 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 2 | 0.034907 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.999390827019096 | 0.999723059691881 | 0.999958776817741+3.3878607017316E-07i | 0.194362848073965 | -0.194682651051888 | 0.44307171379996+0.00165104024854104i | -0.443261168185111-0.00165096752140052i | 21.1942801762784 | 0.219638746278148+0.424985144930845i | -0.220042232009431-0.425102764196957i | -0.999394301906052-0.000635626473230322i | 0.7851 | -3.141 | 44.9826 | 179.9636
|  | 44.9611 | 179.9173 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 3 | 0.05236 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.998629534754574 | 0.99937709285381 | 0.999907292537673+7.61920933088532E-07i | 0.194162734142064 | -0.194882722920739 | 0.442953308197507+0.00165108568816661i | -0.443379533592346-0.00165092188973646i | 21.1869456269482 | 0.21320258372971+0.425770305701788i | -0.214107738306499-0.426040347831904i | -0.998641539283065-0.0014418766396986i | 0.78472 | -3.1401 | 44.9611 | 179.9173
|  | 44.9312 | 179.8511 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 4 | 0.069813 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.997564050259824 | 0.998893128615772 | 0.999835297938027+1.35366090439373E-06i | 0.193882143524031 | -0.195163200046858 | 0.442787558392188+0.00165114933096975i | -0.443545175222875-0.00165085773027746i | 21.1766855118525 | 0.204155706479302+0.426715292375368i | -0.205757722329524-0.427208759697706i | -0.997596370187332-0.00259229341511594i | 0.7842 | -3.139 | 44.9312 | 179.8511
|  | 44.8932 | 179.7637 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 5 | 0.087266 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.996194698091746 | 0.998271556419454 | 0.999742876313662+2.11335770923438E-06i | 0.193520643657134 | -0.195524459917097 | 0.44257448279813+0.00165123120393253i | -0.443758022844215-0.00165077476895198i | 21.1635080872139 | 0.192467347125313+0.427663300649232i | -0.194954901576258-0.428461079919689i | -0.996268820494981-0.00410805741821532i | 0.78353 | -3.1375 | 44.8932 | 179.7637
|  | 44.8475 | 179.6535 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 6 | 0.10472 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.994521895368273 | 0.997512877127242 | 0.999630134618588+3.04017865483136E-06i | 0.193077676312162 | -0.195966989269959 | 0.442314105628812+0.0016513313412528i | -0.444017985739124-0.00165067265489666i | 21.1474239713909 | 0.178106041055779+0.428409311555893i | -0.181657925212714-0.429604659959294i | -0.994673853653723-0.00601564113046618i | 0.78274 | -3.1355 | 44.8475 | 179.6535
|  | 44.795 | 179.5184 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 7 | 0.12217 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.992546151641322 | 0.996617703135246 | 0.999497203361471+4.13310723462829E-06i | 0.192552556023954 | -0.196491385148412 | 0.442006457302428+0.00165144978397324i | -0.444324952391951-0.00165055096159792i | 21.1284461472735 | 0.16104614211602+0.428698006909929i | -0.165826732825863-0.430398555743391i | -0.992833255107571-0.00834610797689743i | 0.78182 | -3.1332 | 44.795 | 179.5184
|  | 44.7364 | 179.3562 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 8 | 0.13963 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.990268068741571 | 0.99558675851062 | 0.999344236477931+5.39094386872408E-06i | 0.191944468160187 | -0.19709835619126 | 0.441651574939259+0.00165158657952748i | -0.444678790105423-0.00165040918829382i | 21.1065899651952 | 0.141276251174086+0.428221818542672i | -0.147429578677853-0.430550718019171i | -0.990778343186152-0.0111340716072761i | 0.7808 | -3.1304 | 44.7364 | 179.3562
|  | 44.6733 | 179.1645 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 9 | 0.15708 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.987688340595138 | 0.994420879152724 | 0.999171411179704+6.81230678075667E-06i | 0.191252466620475 | -0.197788724165776 | 0.441249502952218+0.0016517417812024i | -0.445079344547965-0.00165024676163945i | 21.0818731463487 | 0.118809782055425+0.426619564789467i | -0.126452008979967-0.429715216859984i | -0.988553811136892-0.014416136791489i | 0.7797 | -3.127 | 44.6733 | 179.1645
|  | 44.6075 | 178.9411 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 10 | 0.17453 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.984807753012208 | 0.993121012977465 | 0.998978927780664+8.39563301349642E-06i | 0.190475471156411 | -0.19856342574483 | 0.44080029373184+0.0016519154475168i | -0.445526439231692-0.00165006303764094i | 21.054315786695 | 0.0936978597203505+0.423476292699384i | -0.102908121139067-0.427489908926063i | -0.986223115910673-0.0182285195491736i | 0.77855 | -3.1231 | 44.6075 | 178.9411
|  | 44.5419 | 178.6839 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 11 | 0.19199 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.981627183447664 | 0.991688220124086 | 0.998767009499738+0.0000101391795836699i | 0.189612264301514 | -0.19942351453216 | 0.440304008427234+0.00165210764151521i | -0.446019874921245-0.00164985730386242i | 21.0239403613462 | 0.0660446559149294+0.418325140863275i | -0.0768544512057132-0.423415136150828i | -0.983875962204159-0.0226033280287719i | 0.7774 | -3.1186 | 44.5419 | 178.6839
|  | 44.4804 | 178.3917 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 12 | 0.20944 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.978147600733806 | 0.990123673183581 | 0.998535902240797+0.0000120410247778811i | 0.188661487898371 | -0.200370163339989 | 0.439760717824705+0.00165231842997701i | -0.44655942897373-0.00164962878191158i | 20.9907717294091 | 0.0360250903199809+0.410652255165441i | -0.0484068064533448-0.416974282565406i | -0.981638567533262-0.0275626059118828i | 0.77633 | -3.1135 | 44.4804 | 178.3917
|  | 44.4285 | 178.0644 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 13 | 0.22689 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.974370064785235 | 0.988428657447701 | 0.998285874349506+0.0000140990695914276i | 0.187621639208714 | -0.201404666723553 | 0.4391705033258+0.00165254788253994i | -0.44714485461001-0.00164937663021005i | 20.9548371392632 | 0.00390452847216455+0.399906003082678i | -0.0177602498296075-0.407597316810392i | -0.979687476023531-0.0331085738310408i | 0.77542 | -3.1078 | 44.4285 | 178.0644
|  | 44.3936 | 177.7049 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 14 | 0.24435 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.970295726275996 | 0.986604571177363 | 0.998017216347244+0.0000163110393113481i | 0.186491066590175 | -0.202528443777554 | 0.438533458026785+0.00165279607073795i | -0.44777588011781-0.00164909994705426i | 20.9161662342516 | -0.0299403395921044+0.385511894903782i | 0.0147877968310407-0.394668804756085i | -0.978267551351736-0.039207349579804i | 0.77481 | -3.1015 | 44.3936 | 177.7049
|  | 44.3857 | 177.3203 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 15 | 0.2618 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.965925826289068 | 0.984652925889157 | 0.99773024064214+0.0000186744852461251i | 0.185267964721773 | -0.203743041200038 | 0.437849687901576+0.00165306306695269i | -0.448452207987051-0.00164879777397412i | 20.8747910587548 | -0.0650008733850863+0.36689465817686i | 0.0488117973814804-0.377542260142838i | -0.97771401390717-0.0457614865013498i | 0.77468 | -3.0948 | 44.3857 | 177.3203
|  | 44.418 | 176.925 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 16 | 0.27925 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.961261695938319 | 0.982575346658402 | 0.9974252812173+0.0000211867866035349i | 0.183950369358092 | -0.205050136629494 | 0.43711931309032+0.001653348943279i | -0.449173513978053-0.00164846909939581i | 20.8307460646155 | -0.100618319485536+0.343508720302081i | 0.0837308613130524-0.355563043124687i | -0.97847603576695-0.0525636131144098i | 0.77524 | -3.0879 | 44.418 | 176.925
|  | 44.5065 | 176.5459 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 17 | 0.29671 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.956304755963036 | 0.980373572436999 | 0.997102693296285+0.0000238451525194907i | 0.182536151590101 | -0.206451542261408 | 0.436342469295858+0.00165365377030341i | -0.449939446123302-0.00164811286261622i | 20.7840681178772 | -0.135971196193286+0.314877805829167i | 0.118782223649354-0.328102190139163i | -0.981133376340592-0.0592193305589055i | 0.77678 | -3.0813 | 44.5065 | 176.5459
|  | 44.6694 | 176.2284 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 18 | 0.31416 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.951056516295154 | 0.978049456384073 | 0.996762852986001+0.000026646624239256i | 0.181023011588364 | -0.207949208750797 | 0.435519309290253+0.00165397761579691i | -0.450749623663707-0.00164772795809559i | 20.7347965057939 | -0.170072130499513+0.280643323488744i | 0.152999050789255-0.294603359109026i | -0.986383752659877-0.0650242165824023i | 0.77963 | -3.0758 | 44.6694 | 176.2284
|  | 44.9233 | 176.0452 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 19 | 0.33161 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.945518575599317 | 0.975604966207208 | 0.996406156897005+0.000029588077454093i | 0.179408471801973 | -0.209545229407591 | 0.434650004533806+0.0016543205433211i | -0.45160363592036-0.00164731324007592i | 20.6829729440641 | -0.20177857352047+0.240619627708746i | 0.185197873875923-0.254644214594396i | -0.994950979888511-0.0687852557822731i | 0.78406 | -3.0726 | 44.9233 | 176.0452
|  | 45.2752 | 176.1033 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 20 | 0.34907 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.939692620785908 | 0.973042184511706 | 0.996033021741464+0.0000326662247950178i | 0.177689869584193 | -0.211241844691969 | 0.433734746908635+0.00165468261074951i | -0.452501041103112-0.00164686752753153i | 20.6286415842354 | -0.229821697703158+0.194852139806287i | 0.213982687711538-0.208011721129517i | -1.00732029319911-0.0686143737503723i | 0.7902 | -3.0736 | 45.2752 | 176.1033
|  | 45.7086 | 176.5415 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 21 | 0.36652 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.933580426497202 | 0.970363309155017 | 0.995643883908822+0.0000358776184864747i | 0.175864349213073 | -0.213041447016933 | 0.432773750569073+0.00165506386870413i | -0.453441365057421-0.00164638960945848i | 20.5718490212186 | -0.25285663642493+0.143671964316714i | 0.237774004287481-0.154787674978861i | -1.0231774218499-0.0618372472901648i | 0.79776 | -3.0812 | 45.7086 | 176.5415
|  | 46.1654 | 177.5026 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 22 | 0.38397 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.927183854566788 | 0.967570653603136 | 0.995239199019331+0.0000392186531622225i | 0.173928853272452 | -0.2149465858656 | 0.431767253910895+0.00165546435890882i | -0.454424099951265-0.00164587825050895i | 20.5126443008426 | -0.269534787767908+0.0877385607205412i | 0.254871073709126-0.0954363532510674i | -1.04054022215954-0.0453829590195938i | 0.80574 | -3.098 | 46.1654 | 177.5026
|  | 46.5386 | 179.0687 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 23 | 0.40143 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.92050485345244 | 0.964666647285455 | 0.99481944145566+0.0000426855688456264i | 0.171880113355813 | -0.2169599732308 | 0.430715521661248+0.00165588411246146i | -0.45544870290412-0.0016453321969758i | 20.4510789273785 | -0.278595184520744+0.0280609874580767i | 0.263553024051777-0.0308809042475711i | -1.05506482859047-0.0171516472059814i | 0.81225 | -3.1253 | 46.5386 | 179.0687
|  | 46.6949 | 178.8231 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 24 | 0.41888 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.913545457642601 | 0.961653835943972 | 0.994385103872705+0.000046274454096959i | 0.169714640052205 | -0.219084489384464 | 0.429618847090844+0.00165632314802673i | -0.456514594560422-0.00164475018313204i | 20.3872068709419 | -0.278967425005949-0.0340109654408573i | 0.262218906110245+0.0374496303524842i | -1.06076007507556+0.0217913625991309i | 0.81498 | 3.1211 | 46.6949 | 178.8231
46.5798 | 170.754 | 46.5329 | 176.4097 | 1.0178e-06 | 0.0010278 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 25 | 0.43633 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.90630778703665 | 0.958534881972475 | 0.993936696685862+0.0000499812493297462i | 0.167428712169932 | -0.221323188984361 | 0.428477554349816+0.00165678146995254i | -0.457621157610223-0.00164413093792865i | 20.321084574684 | -0.269874397392558-0.0968312609975602i | 0.24955742389532+0.107622807868499i | -1.05292108795985+0.0660652073390125i | 0.81215 | 3.0789 | 46.5329 | 176.4097
|  | 46.0396 | 174.0181 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 26 | 0.45379 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.898794046299167 | 0.955312564740735 | 0.99347474753793+0.0000538017502987063i | 0.165018365150021 | -0.223679307525437 | 0.427292000928099+0.0016572590663135i | -0.458767735260181-0.00164347319205339i | 20.2527709616646 | -0.250920401082388-0.15853778742903i | 0.224725894898171+0.177271765213833i | -1.03131541892448+0.108066395227661i | 0.80354 | 3.0372 | 46.0396 | 174.0181
|  | 45.2913 | 171.9205 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 27 | 0.47124 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.891006524188368 | 0.951989780898233 | 0.99299980074496+0.0000577316117609061i | 0.162479378617442 | -0.226156268142813 | 0.426062580240851+0.0016577559068854i | -0.459953629658476-0.00164277568535106i | 20.1823274412911 | -0.222150546366963-0.217161598976348i | 0.187508044376015+0.24373762793066i | -1.00019213347387+0.141982490405005i | 0.79048 | 3.0006 | 45.2913 | 171.9205
|  | 44.4041 | 170.2511 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 28 | 0.48869 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.882947592858927 | 0.948569544651335 | 0.99251241672123+0.0000617663513127148i | 0.159807263013692 | -0.228757688773017 | 0.424789724338821+0.00165827194105533i | -0.461178100277679-0.00164203717460502i | 20.1098179151948 | -0.184071054557304-0.270752413476834i | 0.138416188733154+0.304276157591251i | -0.96526759814879+0.165844057297882i | 0.775 | 2.9714 | 44.4041 | 170.2511
|  | 43.4818 | 169.0219 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 29 | 0.50615 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.874619707139396 | 0.945054988007208 | 0.992013171383695+0.0000659013534042647i | 0.156997245249776 | -0.231487389679496 | 0.423473906742935+0.0016588070956736i | -0.462440362260223-0.00164125644167701i | 20.0353087824016 | -0.137626694219364-0.31750378320132i | 0.0787090884588906+0.35630132308825i | -0.931005957656753+0.180600126277012i | 0.7589 | 2.95 | 43.4818 | 169.0219
43.028 | 166.6645 | 42.5934 | 168.1831 | 0.0001041 | 8.1524e-05 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 30 | 0.5236 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.866025403784439 | 0.941449360977141 | 0.991502655536282+0.0000701318735323537i | 0.154044253313383 | -0.234349401347734 | 0.422115645401631+0.00165936127285321i | -0.463739584731618-0.00164043230200023i | 19.9588689436427 | -0.0841393838749344-0.355861316607651i | 0.0103128480494282+0.397624949391952i | -0.899850169319498+0.188266269322298i | 0.74339 | 2.9353 | 42.5934 | 168.1831
|  | 41.7738 | 167.6719 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 31 | 0.54105 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.857167300702113 | 0.937756031731072 | 0.990981474234219+0.0000744530426145371i | 0.150942899758734 | -0.237347972754343 | 0.420715505768553+0.001659934347725i | -0.465074889087223-0.00163956361341885i | 19.8805698046306 | -0.0252187709223198-0.384601114246273i | -0.0643454539140098+0.42664883482716i | -0.872682403240522+0.190723399759347i | 0.72909 | 2.9264 | 41.7738 | 167.6719
|  | 41.0343 | 167.437 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 32 | 0.55851 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.848048096156426 | 0.933978486694525 | 0.990450246128956+0.0000788598715445564i | 0.147687464001524 | -0.240487580013469 | 0.419274103997269+0.00166052616615649i | -0.466445347259052-0.0016386492853634i | 19.8004852781141 | 0.0373404698475345-0.402872001948351i | -0.142560939947235+0.442475983280427i | -0.849500958844916+0.189310128677394i | 0.71618 | 2.9223 | 41.0343 | 167.437
|  | 40.3723 | 167.4442 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 33 | 0.57596 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.838670567945424 | 0.930120330578146 | 0.989909602793938+0.000083347255930565i | 0.144271873334971 | -0.243772935402421 | 0.417792110248512+0.00166113654244387i | -0.467849979969846-0.00163768828834906i | 19.7186917845044 | 0.101659562461833-0.410202252101015i | -0.22158881779765+0.444927478977982i | -0.829900274501475+0.184832499992574i | 0.70463 | 2.9225 | 40.3723 | 167.4442
|  | 39.7779 | 167.6754 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 34 | 0.59341 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.829037572555042 | 0.926185286329265 | 0.989360188031778+0.0000879099810167881i | 0.140689682575939 | -0.247208996766896 | 0.416270252104155+0.00166176525698796i | -0.469287754982398-0.00163667966378024i | 19.6352682508484 | 0.165862592133724-0.406477386585117i | -0.298855991213413+0.434473575738686i | -0.813330446529016+0.177700299762737i | 0.69426 | 2.9265 | 39.7779 | 167.6754
38.8613 | 164.7469 | 39.2383 | 168.1245 | 9.2307e-05 | 0.00040362 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 35 | 0.61087 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.819152044288992 | 0.922177194993823 | 0.98880265716318+0.0000925427267888873i | 0.136934052242436 | -0.250800977304235 | 0.414709318080684+0.00166241205396605i | -0.470757585352994-0.00163562253404105i | 19.5502961079015 | 0.228142077839098-0.391899173981774i | -0.372117319288489+0.41210414920319i | -0.799215023327781+0.168063520618848i | 0.68484 | 2.9343 | 39.2383 | 168.1245
|  | 38.7398 | 168.794 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 36 | 0.62832 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.809016994374948 | 0.918100015476224 | 0.988237676298269+0.0000972400732627134i | 0.132997725155469 | -0.254554355721051 | 0.413110161233362+0.00166307663901209i | -0.472258327698696-0.00163451611284718i | 19.4638592850362 | 0.286800894085748-0.366936833204238i | -0.439544329349647+0.379170547656197i | -0.786995053238576+0.155914849285962i | 0.67614 | 2.946 | 38.7398 | 168.794
|  | 38.2689 | 169.6921 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 37 | 0.64577 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.798635510047293 | 0.913957824183453 | 0.987665921590672+0.000101996505957048i | 0.12887300134901 | -0.258474886758972 | 0.411473702840522+0.00166375867691973i | -0.473788780489099-0.00163335971583083i | 19.3760442026975 | 0.34027268838693-0.332280247418944i | -0.499749765223548+0.337228392751841i | -0.776139279025474+0.141158737253343i | 0.66792 | 2.9617 | 38.2689 | 169.6921
|  | 37.8134 | 170.8314 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 38 | 0.66323 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.788010753606722 | 0.909754814538631 | 0.98708807847512+0.0001068064215481i | 0.124551711162005 | -0.262568612079274 | 0.409800936155363+0.00166445778938382i | -0.475347682374362-0.00163215277132434i | 19.2869397620903 | 0.387126014164086-0.288802524376477i | -0.551761604195826+0.287902496729108i | -0.766140558561982+0.123656202610347i | 0.65997 | 2.9816 | 37.8134 | 170.8314
|  | 37.3625 | 172.2278 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 39 | 0.68068 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.777145961456971 | 0.905495296348442 | 0.986504840889042+0.000111664133705699i | 0.120025186375174 | -0.266841871494118 | 0.40809293021063+0.00166517355279674i | -0.476933710562057-0.00163089483130465i | 19.1966373317686 | 0.42605909847277-0.237536309030084i | -0.594964167512534+0.232785006591536i | -0.756507987406724+0.10325427235711i | 0.6521 | 3.0059 | 37.3625 | 172.2278
35.6943 | 169.1093 | 36.907 | 173.8999 | 0.0010797 | 0.0007589 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 40 | 0.69813 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.766044443118979 | 0.90118369500708 | 0.985916910478843+0.00011656387910892i | 0.115284229243531 | -0.271301314528006 | 0.406350833659152+0.00166590549611775i | -0.478545479256809-0.00162958558245398i | 19.1052307307591 | 0.455892773218646-0.179665378939564i | -0.629023013268574+0.17336931522917i | -0.746759322810276+0.0798066611936721i | 0.64415 | 3.0351 | 36.907 | 173.8999
|  | 36.4397 | 175.8688 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 41 | 0.71558 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.754709580222772 | 0.8968245505188 | 0.98532499579159+0.000121499823638876i | 0.110319079262102 | -0.275953912289196 | 0.404575878630706+0.00166665309883395i | -0.480181538177521-0.00162822485728716i | 19.0128162078392 | 0.47556897557716-0.116530420370094i | -0.653806643494801+0.111016646201018i | -0.736416415651744+0.0531899454419342i | 0.63599 | 3.0695 | 36.4397 | 175.8688
|  | 35.9564 | 178.1579 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 42 | 0.73304 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.743144825477394 | 0.892422516319433 | 0.984729811452773+0.000126466068746i | 0.105119377487897 | -0.280806969625828 | 0.402769384583023+0.00166741578903426i | -0.481840371168345-0.00162681264528973i | 18.9194924165527 | 0.484161608161154-0.0496454292083127i | -0.669313670574468+0.0469495689220655i | -0.725005891290638+0.0233174904867313i | 0.62756 | 3.1094 | 35.9564 | 178.1579
|  | 35.456 | 179.2076 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 43 | 0.75049 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.731353701619171 | 0.887982357876278 | 0.984132077330959+0.000131456657988752i | 0.0996741282248422 | -0.285868137536404 | 0.400932762121794+0.00166819294161598i | -0.483520394920618-0.00162534910400489i | 18.8253603855277 | 0.480905452645946+0.0192812644822081i | -0.675610266893207-0.0177336727511116i | -0.712067481586021-0.00984841870099224i | 0.61882 | -3.1278 | 35.456 | 179.2076
|  | 34.9423 | 176.2015 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 44 | 0.76794 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.719339800338652 | 0.883508951044484 | 0.983532517690096+0.000136465583739382i | 0.0939716578607294 | -0.291145425798348 | 0.399067516761632+0.00166898387664782i | -0.485219957824187-0.0016238345700006i | 18.7305234836314 | 0.465247004566945+0.0883298862992858i | -0.672780067664379-0.0820291256392582i | -0.697172766032438-0.0462876582985762i | 0.60986 | -3.0753 | 34.9423 | 176.2015
33.2109 | 172.6212 | 34.4253 | 172.7982 | 0.0012445 | 1.049e-06 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 45 | 0.7854 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.707106781186548 | 0.879007280156736 | 0.982931860330386+0.000141486794052404i | 0.0879995706252211 | -0.296647215771906 | 0.397175252596523+0.00166978785791159i | -0.486937338967805-0.00162226956964168i | 18.635087379471 | 0.436918096637437+0.155367027248361i | -0.660887348293525-0.144984860480444i | -0.67995735009072-0.0859200979704425i | 0.60084 | -3.0159 | 34.4253 | 172.7982
|  | 33.924 | 168.9756 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 46 | 0.80285 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.694658370458998 | 0.874482435822097 | 0.982330835718565+0.000146514199690854i | 0.0817447010160251 | -0.302382273329664 | 0.395257675845221+0.00167060409164608i | -0.488670747309309-0.00162065482958522i | 18.5391599947291 | 0.39602853903778+0.218071176774519i | -0.639954186324584-0.205648234789048i | -0.660169501640351-0.128616213253223i | 0.59209 | -2.9492 | 33.924 | 168.9756
|  | 33.4691 | 164.7197 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 47 | 0.8203 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.681998360062499 | 0.869939612408564 | 0.981730176108522+0.000151541681304415i | 0.0751930626157524 | -0.30835976185414 | 0.393316598233412+0.00167143172551761i | -0.490418321037531-0.00161899128691067i | 18.4428514507935 | 0.343167577887815+0.273998779256161i | -0.609953359756428-0.263003040435036i | -0.637737968890874-0.174229312351742i | 0.58415 | -2.8749 | 33.4691 | 164.7197
|  | 33.1066 | 160.0332 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 48 | 0.83776 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.669130606358859 | 0.865384105182873 | 0.981130614653287+0.000156563096752503i | 0.0683297929945971 | -0.31458925523748 | 0.391353940170983+0.00167226984784151i | -0.492178127148984-0.00161728009878895i | 18.3462740081204 | 0.279496394288774+0.320700287088779i | -0.57082092542314-0.315896168690627i | -0.612861572465876-0.222661186281633i | 0.57782 | -2.7931 | 33.1066 | 160.0332
|  | 32.9027 | 154.9469 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 49 | 0.85521 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.656059028990507 | 0.860821307080526 | 0.980532884509336+0.000161572288566191i | 0.0611390943632045 | -0.321080750808453 | 0.389371733679353+0.00167311748707828i | -0.493948161263169-0.00161552265158962i | 18.2495419977587 | 0.206807008602079+0.355887667553313i | -0.522495834741188-0.362953960365412i | -0.586124719903527-0.273976115555387i | 0.57426 | -2.7043 | 32.9027 | 154.9469
32.2015 | 150.2136 | 32.948 | 149.5372 | 0.00051331 | 2.0464e-05 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 50 | 0.87266 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.642787609686539 | 0.856256705077667 | 0.979937717934327+0.000166563091539945i | 0.053604169605811 | -0.327844682101827 | 0.387372125020053+0.00167397361162848i | -0.495726347701547-0.00161372056931722i | 18.1527717444331 | 0.127519979499777+0.377641923057544i | -0.464998565366536-0.402491230720898i | -0.558649923501516-0.328582148817699i | 0.57505 | -2.6099 | 32.948 | 149.5372
|  | 33.3619 | 143.9469 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 51 | 0.89012 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.629320391049837 | 0.851695876136507 | 0.979345845379335+0.000171529340445418i | 0.0457071532849079 | -0.334891931375208 | 0.385357376972532+0.00167483712994932i | -0.497510539855649-0.00161187572126379i | 18.0560814805867 | 0.0445978181901365+0.384633551335454i | -0.398566028340596-0.432422838606028i | -0.532318905630083-0.387505929098543i | 0.58228 | -2.5123 | 33.3619 | 143.9469
|  | 34.2985 | 138.4079 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 52 | 0.90757 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.615661475325658 | 0.847144482695165 | 0.978757994576721+0.000176464877857981i | 0.0374290371652426 | -0.34223384176714 | 0.383329870705782+0.00167570689101496i | -0.499298520870632-0.00160999022875839i | 17.9595912507644 | -0.0386347413881306+0.376315480472431i | -0.323863883890788-0.450201225473897i | -0.510146712249565-0.452803825126883i | 0.59862 | -2.4157 | 34.2985 | 138.4079
|  | 35.9523 | 133.2591 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 53 | 0.92502 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.601815023152048 | 0.842608267673021 | 0.978174889624838+0.000181363562085658i | 0.0287495897556212 | -0.349882228978806 | 0.381292107185395+0.00167658168514045i | -0.501088004670806-0.00160806647089019i | 17.8634228057264 | -0.118750142632028+0.353043463581077i | -0.242294488884225-0.452826335665366i | -0.497026512723085-0.5281868757539i | 0.62749 | -2.3258 | 35.9523 | 133.2591
|  | 38.568 | 128.9642 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 54 | 0.94248 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.587785252292474 | 0.838093048962717 | 0.97759725007077+0.00018621927519007i | 0.0196472693113906 | -0.357849392349537 | 0.379246708054884+0.00167746024518822i | -0.502876637353917-0.00160610708907833i | 17.7676994856773 | -0.192519586678351+0.316088256218394i | -0.156397993442223-0.437006593866435i | -0.501417723690726-0.619991360182443i | 0.67314 | -2.2508 | 38.568 | 128.9642
41.4376 | 129.2823 | 42.4495 | 126.1505 | 0.0005682 | 0.00061631 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 55 | 0.95993 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.573576436351045 | 0.833604713380627 | 0.977025789992353+0.000191025931087958i | 0.0100991296772309 | -0.366148125183782 | 0.377196415927988+0.00167834124817321i | -0.504661998980765-0.00160411499035908i | 17.6725460920152 | -0.257203067709232+0.267526893876698i | -0.0702885278582081-0.399580430634154i | -0.539595320407907-0.738600806729937i | 0.74088 | -2.2017 | 42.4495 | 126.1505
|  | 47.9394 | 125.7514 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 56 | 0.97738 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.559192903470747 | 0.829149210048618 | 0.976461217080739+0.000195777483721371i | 0.0000807182781223729 | -0.374791724174434 | 0.37514409402681+0.00167922331727923i | -0.506441605786263-0.00160209334926025i | 17.578088747024 | -0.310773523136051+0.210030027728926i | 0.0100372984481408-0.338303212483376i | -0.647518831271986-0.899412633838748i | 0.8367 | -2.1948 | 47.9394 | 125.7514
|  | 55.264 | 129.4964 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 57 | 0.99484 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.544639035015027 | 0.824732543181288 | 0.97590423172491+0.000200467935283204i | -0.01043403451548 | -0.383793997754283 | 0.373092725099685+0.00168010502429769i | -0.508212912837337-0.00160004560813195i | 17.4844547409621 | -0.352037173547464+0.146587469072694i | 0.0768174489806227-0.253000293050192i | -0.917310930979201-1.11293028364889i | 0.96454 | -2.2601 | 55.264 | 129.4964
|  | 63.9047 | 141.4937 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 58 | 1.0123 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.529919264233205 | 0.820360764254902 | 0.975355526099337+0.000205091344485641i | -0.0214729355155213 | -0.393169273194427 | 0.371045409552438+0.00168098489249377i | -0.509973317161676-0.00159797547580679i | 17.391772366042 | -0.380639210803613+0.0802254088910997i | 0.121428221759595-0.146823556697073i | -1.59769077614679-1.27114793520134i | 1.1153 | -2.4695 | 63.9047 | 141.4937
|  | 70.7282 | 170.7809 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 59 | 1.0297 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.515038074910055 | 0.816039963536756 | 0.97481578325631+0.000209641834856406i | -0.0330656593114525 | -0.402932402255625 | 0.36900536272736+0.00168186139990334i | -0.511720161369797-0.00159588692446466i | 17.3001707368524 | -0.396971056165913+0.0137633976139048i | 0.136233997908139-0.0269869288412856i | -2.82312296286937-0.458211766914026i | 1.2344 | -2.9807 | 70.7282 | 170.7809
64.3964 | 147.7017 | 69.5371 | 147.821 | 0.0054653 | 6.5112e-07 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 60 | 1.0472 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.499999999999999 | 0.811776260956676 | 0.974285676224226+0.000214113603048757i | -0.045243890428991 | -0.413098765186076 | 0.36697591126594+0.00168273298305795i | -0.513450737790702-0.00159378418458488i | 17.2097795968316 | -0.402013222842384-0.0503583906942545i | 0.116958929931908+0.0957945692619156i | -2.26824980731351+1.42723281303061i | 1.2137 | 2.58 | 69.5371 | 147.821
|  | 61.9976 | 122.6156 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 61 | 1.0647 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.484809620246337 | 0.807575796306305 | 0.973765867113337+0.000218500927149879i | -0.0580414782896726 | -0.423684272847451 | 0.364960488494456+0.001683598041131i | -0.515162293138661-0.00159167173787636i | 17.120729110492 | -0.397152700053746-0.110172445823987i | 0.0644851607673305+0.209464965954224i | -1.01361153886204+1.5839870667781i | 1.0821 | 2.14 | 61.9976 | 122.6156
|  | 53.9468 | 110.3123 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 62 | 1.0821 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.46947156278589 | 0.803444718756075 | 0.973257006230308+0.00022279817497194i | -0.071494606370044 | -0.434705366740206 | 0.362962628775762+0.00168445494049331i | -0.516852033725491-0.00158955430808768i | 17.0331496411773 | -0.384007833468743-0.164206221001477i | -0.015062593056848+0.303654840869024i | -0.476862067328127+1.28827392184367i | 0.94155 | 1.9253 | 53.9468 | 110.3123
|  | 47.3029 | 103.3618 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 63 | 1.0996 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.453990499739546 | 0.799389175685037 | 0.972759731203139+0.000226999812309099i | -0.0856419771458465 | -0.446179016689644 | 0.360985960776359+0.00168530201966098i | -0.518517131228825-0.00158743684961239i | 16.9471715142538 | -0.364282103599886-0.211451985811318i | -0.112145194349322+0.372104921660687i | -0.250464669072228+1.05446203409377i | 0.82559 | 1.804 | 47.3029 | 103.3618
|  | 42.1573 | 98.85 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 64 | 1.117 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.438371146789078 | 0.79541530082439 | 0.972274666117841+0.000231100411144382i | -0.100525014621996 | -0.458122715946446 | 0.359034199605106+0.00168613759461327i | -0.520154729022496-0.00158532453382171i | 16.8629247657513 | -0.33965612061517-0.251305360399019i | -0.216546004028992+0.413386618939771i | -0.139291966506797+0.894608636109333i | 0.73578 | 1.7253 | 42.1573 | 98.85
39.6924 | 97.2101 | 38.1828 | 95.6566 | 0.0015633 | 0.00026372 | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 65 | 1.1345 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.422618261740699 | 0.791529201722272 | 0.971802420668373+0.000235094657788601i | -0.116188086498588 | -0.47055447344945 | 0.35711113778893+0.00168695996445314i | -0.521761949070118-0.00158322273307569i | 16.7805388766147 | -0.311717763079214-0.28349106345712i | -0.319651878654974+0.429968647836752i | -0.0775163651498117+0.78260593305774i | 0.66642 | 1.6695 | 38.1828 | 95.6566
|  | 35.0376 | 93.2764 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 66 | 1.1519 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.4067366430758 | 0.787736946544545 | 0.971343589321338+0.000238977360934178i | -0.132678748312312 | -0.483492802995112 | 0.355220635061413+0.00168776741737808i | -0.52333589937735-0.00158113700238609i | 16.7001424928787 | -0.281924728449348-0.307989406101382i | -0.415543286973802+0.426515499684059i | -0.0400751831170583+0.700039510847926i | 0.61152 | 1.628 | 35.0376 | 93.2764
|  | 32.4519 | 91.4333 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 67 | 1.1694 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.390731128489275 | 0.784044550234302 | 0.970898750496867+0.00024274345960624i | -0.150048012228878 | -0.496956709057465 | 0.353366606952519+0.00168855823692362i | -0.524873681992107-0.00157907305872948i | 16.6218631322473 | -0.251591402658068-0.324971790268645i | -0.500992911788214+0.408311029793369i | -0.0159058399210497+0.635692148326009i | 0.56639 | 1.5958 | 32.4519 | 91.4333
|  | 30.2288 | 89.9427 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 68 | 1.1868 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.374606593415913 | 0.780457960061322 | 0.970468465767234+0.000246388030992214i | -0.168350643553271 | -0.5109656690049 | 0.35155301218166+0.00168933070843875i | -0.526372401535351-0.00157703675803562i | 16.5458268777399 | -0.221891853405073-0.33474909126788i | -0.574872932485267+0.380234109797456i | 0.000582241321346306+0.582686121314251i | 0.52759 | 1.5698 | 30.2288 | 89.9427
|  | 28.228 | 88.6562 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 69 | 1.2043 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.358367949545299 | 0.776983040601787 | 0.970053279074608+0.000249906298132617i | -0.187645488481392 | -0.525539611466394 | 0.349783838871831+0.00169008312574719i | -0.527829174237851-0.00157503406990734i | 16.4721580592593 | -0.193871865125465-0.337733691383729i | -0.637427758074576+0.346291109129458i | 0.0125897661145476+0.53667793898488i | 0.49267 | 1.5473 | 28.228 | 88.6562
|  | 26.3514 | 87.4339 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 70 | 1.2217 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.34202014332567 | 0.773625558198382 | 0.96965371596947+0.000253293637454056i | -0.207995837150851 | -0.540698890610046 | 0.348063089619722+0.00169081379794663i | -0.529241137450642-0.0015730710501591i | 16.400978924143 | -0.168464499825777-0.334413985165586i | -0.689659791264865+0.309526635092963i | 0.0221776820938769+0.494850610113642i | 0.45992 | 1.526 | 26.3514 | 87.4339
|  | 24.5317 | 86.1273 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 71 | 1.2392 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.325568154457157 | 0.77039116496031 | 0.969270282871048+0.000256545586126563i | -0.229469826674393 | -0.556464256111184 | 0.346394765474895+0.00169152105629458i | -0.53060545958919-0.0015711538112959i | 16.3324092979616 | -0.146505246518657-0.325339108866914i | -0.732900389379871+0.272127371240306i | 0.0308249185489858+0.455351665459917i | 0.42816 | 1.5032 | 24.5317 | 86.1273
|  | 22.7268 | 84.5618 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 72 | 1.2566 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.309016994374946 | 0.767285382373398 | 0.968903466351285+0.000259657849226684i | -0.252140889575112 | -0.572856818606602 | 0.344782848900546+0.00169220326112849i | -0.531919350462953-0.00156928849108878i | 16.2665662370504 | -0.128744225976104-0.311111157983122i | -0.76854772433374+0.23559293412737i | 0.0396962189944456+0.416972552963576i | 0.39666 | 1.4759 | 22.7268 | 84.5618
|  | 20.9177 | 82.517 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 73 | 1.2741 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.292371704722738 | 0.764313584600054 | 0.968553732443658+0.000262626306688435i | -0.276088253908806 | -0.589898010455377 | 0.343231285808431+0.00169285880876714i | -0.533180071934082-0.0015674812194391i | 16.2035636744658 | -0.11585411604317-0.292382023400726i | -0.797925366128657+0.200907034749457i | 0.0497768347548553+0.378960931067558i | 0.36508 | 1.4402 | 20.9177 | 82.517
|  | 19.1106 | 79.7021 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 74 | 1.2915 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.275637355816999 | 0.76148098155909 | 0.968221525978264+0.000265447020024087i | -0.301397502381874 | -0.607609541655798 | 0.3417439667809+0.00169348613833972i | -0.534384948840814-0.00156573808375921i | 16.1435120612752 | -0.108433494782618-0.269852068889811i | -0.822217634835919+0.168681782667602i | 0.0619405103897607+0.340907677877322i | 0.33354 | 1.3911 | 19.1106 | 79.7021
|  | 17.3442 | 75.7302 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 75 | 1.309 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.258819045102521 | 0.758792601884727 | 0.967907269944449+0.000268116238796961i | -0.328161198984553 | -0.626013350801763 | 0.340324707613472+0.00169408373849158i | -0.535531380113406-0.00156406509313213i | 16.0865180052851 | -0.107006210522378-0.244268142706369i | -0.842449228836047+0.139269025053307i | 0.0769812622306765+0.30267610120186i | 0.30271 | 1.3217 | 17.3442 | 75.7302
|  | 15.7016 | 70.1135 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 76 | 1.3265 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.241921895599668 | 0.756253275873243 | 0.967611364882332+0.00027063040682951i | -0.356479593098623 | -0.645131551001241 | 0.338977229330951+0.00169465015391586i | -0.536616850003041-0.00156246814154631i | 16.0326839095077 | -0.112018194609236-0.216418866485978i | -0.859488025704336+0.112841761894838i | 0.0956243670092392+0.2643542222164i | 0.27404 | 1.2237 | 15.7016 | 70.1135
|  | 14.3264 | 62.348 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 77 | 1.3439 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.224951054343866 | 0.753867618533392 | 0.967334188304414+0.00027298616812923i | -0.386461412752751 | -0.66498637072275 | 0.337705137848932+0.00169518399166491i | -0.537638939337588-0.00156095297053192i | 15.9821076128271 | -0.123833802489426-0.187125789365369i | -0.874058482688829+0.0894515356211098i | 0.118525504556703+0.226218361441869i | 0.25004 | 1.0882 | 14.3264 | 62.348
|  | 13.428 | 52.2203 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 78 | 1.3614 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.207911690817758 | 0.751640012864221 | 0.967076094148472+0.000275180372516659i | -0.418224760748756 | -0.685600089584265 | 0.336511903469661+0.00169568392719734i | -0.538595336711997-0.00155952513155349i | 15.9348820354865 | -0.142734229479868-0.1572298207395i | -0.88675867476558+0.0690677324304661i | 0.146264260343754+0.188700676180357i | 0.23436 | 0.91142 | 13.428 | 52.2203
|  | 13.2418 | 40.3497 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 79 | 1.3788 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.190808995376544 | 0.749574593489002 | 0.966837412262924+0.000277210080939415i | -0.451898129840201 | -0.706994969150756 | 0.335400840415876+0.00169614871012267i | -0.539483849516567-0.00155818994853638i | 15.891094832125 | -0.168920714537329-0.127573290443466i | -0.898077422130599+0.0516037674367005i | 0.179337049597329+0.152356305227071i | 0.23111 | 0.70423 | 13.2418 | 40.3497
|  | 13.9251 | 28.3741 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 80 | 1.3963 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.173648177666931 | 0.747675230778503 | 0.966618447925689+0.000279072570457602i | -0.487621556105294 | -0.729193178862238 | 0.334375086618801+0.00169657716960767i | -0.540302414702438-0.00155695248092334i | 15.8508280551887 | -0.202525064259801-0.0989788931480195i | -0.908409922309096+0.0369339821667524i | 0.218153940856143+0.117828068893315i | 0.24304 | 0.49522 | 13.9251 | 28.3741
|  | 15.466 | 18.0671 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 81 | 1.4137 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.15643446504023 | 0.745945515598955 | 0.966419481397611+0.000280765338886076i | -0.525547933233803 | -0.752216717271971 | 0.333437583985097+0.00169696821941592i | -0.541049109181565-0.0015558174876701i | 15.8141578315852 | -0.24362952100636-0.0722274685937939i | -0.918071375554509+0.0249040781939996i | 0.263043316731076+0.085808491605015i | 0.26993 | 0.31533 | 15.466 | 18.0671
|  | 17.7132 | 10.2803 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 82 | 1.4312 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.139173100960067 | 0.744388744820472 | 0.966240767511357+0.000282286109080613i | -0.565844514790106 | -0.776087328834175 | 0.332591059373092+0.00169732086255572i | -0.541722159757669-0.00155478939259469i | 15.7811540554598 | -0.292297286980591-0.0480368595912899i | -0.927308644524281+0.0153371284184759i | 0.314267605059312+0.0570002474580268i | 0.30915 | 0.17942 | 17.7132 | 10.2803
|  | 20.4773 | 4.9276 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 83 | 1.4486 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.121869343405148 | 0.743007907720034 | 0.966082535296727+0.000283632832855052i | -0.60869463680726 | -0.800826416541113 | 0.331838006509588+0.00169763419551624i | -0.5423199524856-0.00155387225149463i | 15.7518800999371 | -0.348614322556016-0.0270438297525998i | -0.936310244971444+0.00803663696556418i | 0.372052462796174+0.0320768468459434i | 0.3574 | 0.086003 | 20.4773 | 4.9276
|  | 23.595 | 1.5278 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 84 | 1.4661 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.104528463267657 | 0.741805673409427 | 0.965944987643135+0.000284803694518349i | -0.654299699547123 | -0.826454950769102 | 0.331180669075674+0.00169790741207746i | -0.542841041359097-0.00155306972143767i | 15.7263925505908 | -0.412742644131575-0.00979014564777335i | -0.945215053209006+0.00278771857917118i | 0.436630985511537+0.011645339249471i | 0.41181 | 0.026665 | 23.595 | 1.5278
|  | 26.9451 | 0.42699 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 85 | 1.4835 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.0871557427476529 | 0.74078437941209 | 0.965828301000042+0.000285797114020667i | -0.702881455233812 | -0.85299337475167 | 0.330621025181817+0.00169813980668279i | -0.543284156231547-0.00155238503261558i | 15.7047409632716 | -0.484985538186085+0.00328756721527918i | -0.954120110525084-0.000642806860911686i | 0.508304036164139-0.00378810644203434i | 0.47028 | -0.0074523 | 26.9451 | 0.42699
|  | 30.442 | 1.3505 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 86 | 1.501 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.0697564737441267 | 0.739946021505131 | 0.96573262511597+0.000286611749699694i | -0.754684658423582 | -0.880461507153291 | 0.330160773440716+0.00169833077736675i | -0.543648209880213-0.0015518209631275i | 15.6869676487563 | -0.565866118031331+0.0118667080755112i | -0.963087843929917-0.00250680021388673i | 0.587517951254896-0.0138507596014782i | 0.53131 | -0.023571 | 30.442 | 1.3505
|  | 34.0277 | 1.5563 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 87 | 1.5184 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.0523359562429409 | 0.739292244931692 | 0.965658082816684+0.000287246500618319i | -0.809980147924393 | -0.908878442269567 | 0.329801320829029+0.00169847982823399i | -0.543932304132603-0.00155137981702731i | 15.6731074864486 | -0.656222640148702+0.015751752378589i | -0.972152929101532-0.00307520553012312i | 0.674961960409233-0.0183380603999603i | 0.59389 | -0.027162 | 34.0277 | 1.5563
|  | 37.6639 | 1.2812 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 88 | 1.5359 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.0348994967024886 | 0.73882433707708 | 0.96560476982301+0.000287700508487341i | -0.869068444484432 | -0.938262448424926 | 0.329543772508404+0.00169858657148713i | -0.544135734982815-0.00155106340593258i | 15.6631877691117 | -0.757327062719601+0.0149011540585449i | -0.981328910473436-0.00263812173833844i | 0.771689820537747-0.0172592140806174i | 0.65736 | -0.022362 | 37.6639 | 1.2812
|  | 41.3265 | 0.71087 |  |  | 1 | 0 | 1.483 | 0 | 14857 | 3.8435 | 0.015793 | 89 | 1.5533 | 6530 | 1 | 1.483 | 3.8435+0.015793i | 0.0174524064373069 | 0.73854322168733 | 0.965572754608758+0.000287973159166639i | -0.93228396770981 | -0.968630865180854 | 0.329388923750848+0.0016986507290031i | -0.544257996636886-0.00155087303444703i | 15.6572280803014 | -0.871037477466272+0.00948467894282154i | -0.990614578585783-0.00150430744971477i | 0.879273417379499-0.0109097692771193i | 0.72128 | -0.012407 | 41.3265 | 0.71087
Default Solver Settings
0.0069029 | 3 | TRUE | TRUE | TRUE | TRUE
TRUE | TRUE | TRUE | 32767 | 0 | 0
Input known optical system properties here (except for 1 - Film). A compendium of refractive indicies for materials can be accessed through the refractiveindex.info link below this table. Be aware that sufficiently thin materials have different optical properties than bulk.
Solved/fit film parameters.
Input a range of refractive index, extinction coefficient, and film thickness values for the solver to use. Prior knowledge of the substrate and nature of the film can help narrow this range which increases solver speed and accuracy.
Load the default solver settings if they aren't automatically present when using the solver.
The average sum of normalized squared errors should be minimized when using the solver, indicating the experimental data closely matches the fitted modeled values, and the thin film properties (n1, k1, and d1) can be considered solved.